Length of sedimentation reaction in blood: a comparison between plasma citrate and K2-EDTA blood samples.

Zenobi R°, Viola FG°, Ballerini S°, Bellincampi L°, Federici G°

°Department of Laboratory Medicine, University-Hospital of “Tor Vergata” Rome, Italy

The erythrocyte sedimentation rate (ESR), now more appropriately referred to as  "length of sedimentation reaction in blood (LSRB)", remains the most widely used laboratory test for monitoring the course of infections, inflammatory diseases and some types of cancer. In addition, it is simple and low-cost. The most commonly used procedures for LSRB determination requires plasma-citrate blood samples, following ICSH (International Committee for Standardization in Haematology) recommendations. It is no unlikely that LSRB blood samples are insufficient, especially in children and in haemathological patients. In order to resolve this inconvenient, we compared LSRB measurement in plasma citrate and in K2-EDTA blood samples.

We performed our study on 300 whole blood samples from patients of Medicine Laboratory of University Hospital “Tor Vergata” Rome, using an automated procedure (VES-MATIC 60 DIESSE s.p.a.,Grosseto, Italy)
Plasma citrate blood samples were analyzed  and used as control, whereas K2-EDTA blood samples were transferred in tubes containing sodium-citrate and then analyzed.

Our data showed that LRSB values obtained by the two different anticoagulants were the same with a statistically significant correlation (r=0.99). 
In conclusion, LSRB can be eventually measured from the tube collected for hemochrome analysis and so resolve the inconvenient of insufficient samples.
Ozdem S, Akbas HS, Donmez L, Gultekin M.
Comparison of TEST 1 with SRS 100 and ICSH reference method for the measurement of the length of sedimentation reaction in blood.
Clin Chem Lab Med. 2006;44(4):407-12
